# HV Test Program Settings

# (11/21/00)

- For cable continuity self test (insert Cu pieces in connector)

  - HV setting: 500 V

  - Trip Limit: 1

  - Trip Wait V: -1

  - HV Ramp: 3000 V/sec

  - HV stp:   50 V/step

  - Relay duration: 1 sec

  - Data interval: 100 ms

  - Record Data: ON

  - Process Control: ON

  - Module file/ ID : type in proper file name

  - Passing I Draw: 1 uA

  - Quick on I: OFF

  - Relay Mode: SCAN

  - Repeat SCAN: OFF

  - Sequence # Assignment: 1 (initial default)

  - Scanning Sequence Parameter File: select according to HV plate type

- For Cable insulation self test (insert kapton piece in connector),  HV Kapton test, HV plate test

  - HV setting: 3000 V

  - Trip Limit: 3

  - Trip Wait V: -1

  - HV Ramp: 3000 V/sec

  - HV stp:   50 V/step

  - Relay duration: 3 sec

  - Data interval: 100 ms

  - Record Data: ON

  - Process Control: ON

  - Module file/ ID : type in proper file name

  - Passing I Draw: 1 uA

  - Quick on I: OFF

  - Relay Mode: SCAN

  - Repeat SCAN: OFF

  - Sequence # Assignment: 1 (initial default)

  - Scanning Sequence Parameter File: select according to HV plate type

- For Module HV scan after straw insertion, before stringing.

  - HV setting: 2500 V

  - Trip Limit: 5

  - Trip Wait V: -1

  - HV Ramp: 500 V/sec

  - HV stp:   10 V/step

  - Relay duration: 5 sec

  - Data interval: 100 ms

  - Record Data: ON

  - Process Control: ON

  - Module file/ ID : type in proper file name

  - Passing I Draw: 1 uA

  - Quick on I: OFF

  - Relay Mode: SCAN

  - Repeat SCAN: OFF

  - Sequence # Assignment: 1 (initial default)

  - Scanning Sequence Parameter File: select according to HV plate type

- To inspect a particular trace repeatedly (For trouble shooting when tests from above indicate problems at certain trace)

  - Relay duration: set to time desired

  - Record Data: OFF

  - Process Control: ON

  - Relay Mode: Repeat One

  - Sequence # Assignment: type in the HV island number desired

- To set I offset (Usually done if tester not used for a while, in order to properly use “passing I” and “quit on I” option)

  - Relay duration: set to a long time , e.g. 100 sec

  - Run HV tester

  - While running the tester, look at the dial on HV power supply (turn knob to proper range) to read actual current draw

  - Compare actual current draw with the display on screen and type in the difference  at “I Offset”

- To bring the whole module on HV for type 1 or type 2 module

  - Scanning Sequence Parameter File: Connector 1 & 2 ……

  - Relay duration: set to time desired

  - Record Data: OFF

  - Process Control: ON

  - Relay Mode: Repeat One

  - Sequence # Assignment: 73 (or any number > 73)

- To bring the whole module on HV for type 3 module

  - For a type 3 module :

  - Scanning Sequence Parameter File: All 3 Connectors, …..

  - Relay duration: set to time desired

  - Record Data: OFF

  - Process Control: ON

  - Relay Mode: Repeat One

  - Sequence # Assignment: 109 (or any number > 109)