# Scope

## Scope

This procedure establishes requirements for preliminary inspection of the module assembly’s straw-high voltage capacity.

## Applicability

### Applicability

This procedure applies to all modules assembled by the Indiana and Duke Production Facilities.

### Relation to Other ATLAS Project Requirements

The inspection procedure described by this specification is in addition to other tests and inspections required for module assembly. Module-assembly straw-sealing glue application and wire stringing may begin only after acceptable results from this inspection.

# Applicable Documents

## Document List

The following documents of the issue in effect on the effective date of this specification form a part of this specification to the extent specified herein.

### ATL-IT-EY-0004, ATLAS U.S. Environmental, Health, and Safety Plan

## Amendments and Revisions

Whenever this procedure is amended or revised subsequent to its effective date, the Revised Version will be placed in the Engineering Data Management System, the Production Database displays, and released to the technicians.  The QA Engineer and the affected Production Engineers will coordinate release to the technicians.

# Requirements

## Background

High voltage plates distribute high voltage from the fuse box to the straw cathodes.  The high voltage is negative to attract ions and repel electrons.  A Kapton flex circuit on the high voltage plates, which plugs into the fuse box, carries the high voltage on copper traces embedded in the Kapton – first to a resistor and then from the resistor to a feed-through pad.  A thick plate with a collar that mates to the tension plate and the shell carries the mounting holes for mounting the high voltage module in the detector and for supporting and spacing the straws.  A thin plate on the side of the of the high voltage plate opposite the Kapton flex circuit connects each group of eight or seven straws to a feed-through hole.  A pin socket conducts the high voltage through the feed-through hole from the Kapton to the straws.

Conductive glue applied to the inner and outer surfaces of the straws on the thin plate side of the high voltage plates connects the straws to the high voltage circuit.  This test confirms satisfactory straw/glue connections prior to the automatic gluing procedure.

The technician will test all high voltage groups for isolation between groups and for isolation between the shell and the group.  The Kapton HV strips shall be connected to a high voltage test fixture.  The circuits will be tested by sequentially raising each trace and each corresponding group of straws to at least 2.1 kV while holding all the other traces and groups at ground

Appendix A provides a checklist to be used by Technicians performing this procedure.

# Preparation for Delivery

### Storage, Packing , and Shipping Requirements

There are no storage, packing, and shipping requirements applicable to this procedure.

# Environment, Health, and Safety (EH&S)

### EH&S Invoked

This test requires use of high voltage equipment and test methods.  Technicians performing this test must follow the ATLAS U.S. EH&S procedures for high voltage testing as described by reference 2.1.1 above.

# References

## ATLITB1_0010, Module 1 Thick High Voltage Plate

## ATLITB2)0010, Module 2 Thick High Voltage Plate

## ATLITB3_0010, Module 3 Thick High Voltage Plate

## ATLITB1_0001, Type 1 Isometric Assembly Diagram (Exploded)

## ATLITB2_0001, Type 2 Isometric Assembly Diagram (Exploded)

## ATLITB3_0002, Type 3 Isometric Assembly Diagram (Exploded)

Appendix A

Table Name:	tblChecklistVerifyGrpHVStandoff1

Step	Check List Steps

1	Scan the 'Quality Assurance Procedure for Verifying Group High Voltage Standoff # 1" Product ID barcode.

2	Move HV tester to a position close enough to the module end for the cables to reach.  Enter "Done".

3	Attach the set of cables for the type of module to be tested to the resistor cards on the top of the relay box.  Each

cable has a label specifying its postion and orientation.  Enter "Done".

4	Review this checklist and the HV tester instruction booklet.  Logon to ATLASPC24 and double click on hvtest.exe.  Set the sequence file name in the Lab View control panel to "short side acute corner upper left".  Enter "Done".

5	Turn on the control box (greeen lighted switch) and the Bertan HV supply side B.  Also set the trip reset on the Bertan to automatic reset.  Enter "Done".

6	Connect the kapton HV strip cables 1 and 2 to the HV test cable connector 1 and 2.  The "visible side" of each connector should be visible.  Enter "Done".

7	Place the grounding plate or grounding pins against the HV plate and fasten it there. Attach grounding clip to table from the grounding plate to HV Tester.   Enter "Done".

8	Enter module file ID (For example, M2_OIC).  Set to Record Data.  Enter "Done".

9	Run the HV test program with HV=400V and trip level = 1.  If the power supply trips there is a good connection.  Note each connection that doesn't trip.  Enter "Done".

10	Fix the bad connections and enter "Accept".

11	Remove the grounding plate/grounding pins from the HV plate. Connect copper tape between shell and frame.  Enter "Done".

12	In the same HV test program, change to at least HV=2100V, number of trips = 5, and set relay duration time = 15 sec.  Set the button to record data and type a file prefix including the module number.  Enter "Done".

13	Run the HV test program which records current draw data and identifies the groups that trip continuously.  Enter

"Done".

14	Examine the list of files that are preceded by "F". These files are for groups that tripped more than 5 times. The group number is in the file name. Retest these groups and try to locate the discharge. Enter "Done"

15	If the Module dose not trip continuously on any group, enter 'Accept". If the module has HV trip problems, enter "Reject" and return for rework.

16	TURN OFF the Bertan supply and the control box.  Disconnect the HV cables from the kapton strips.  Exit the HV test program.  Enter "Done".

17	Enter "Save" to save the current page and exit the form.  Otherwise enter "Cancel" to exit the form without saving the current page.

Quality Assurance “Procedure to Verify Group High Voltage Standoff #1” Product ID Barcode

Procedure Part Number

# Text boxes

Approvals

| Name | Signature | Revision | Date |
| --- | --- | --- | --- |
| J.Callahan |  | B |  |
| D.Rust |  | B |  |
| C. Wang |  | B |  |